RedPitaya

FPGA memory map

Written by	Revision Description	Version	Date
Matej Oblak	Initial	0.1	08/11/13
Matej Oblak	Release1 update	0.2	16/12/13
Matej Oblak	ASG - added burst mode / ASG - buffer read pointer readout		Dec. 2014 Jan. 2015
Matej Oblak	AXI master documented		Feb. 2015
Iztok Jeras	AVG – added scope averaging		Feb. 2015

# About Document
Red Pitaya HDL design has multiple functions, which are configured by registers. It also uses memory locations to store capture data and generate output signals. All of this are described in this document. Memory location is written in a way that is seen by SW.

# FPGA Memory Map
The table describes address space partitioning implemented on FPGA via AXI GP0 interface. All registers have offsets aligned to 4 bytes and are 32 -bit wide. Granularity is 32-bit, meaning that minimum transfer size is 4 bytes. The organization is little-endian .
The memory block is divided into 8 parts. Each part is occupied by individual IP core. Address space of individual application is described in the subsection below. The size of each IP core address space is 4MByte.
For additional informations and better understanding check other documents (schematics, specifications...).

	Start	End	Module Name
CS[0]	0x40000000	0x400FFFFF	Housekeeping
CS[1]	0x40100000	0x401FFFFF	Oscilloscope
CS[2]	0x40200000	0x402FFFFF	Arbitrary signal generator (ASG)
CS[3]	0x40300000	0x403FFFFF	PID controller
CS[4]	0x40400000	0x404FFFFF	Analog mixed signals (AMS)
CS[5]	0x40500000	0x405FFFFF	Daisy chain
CS[6]	0x40600000	0x406FFFFF	FREE
CS[7]	0x40700000	0x407FFFFF	Power test

# Red Pitaya Modules
Here are described submodules used in Red Pitaya FPGA logic.
## Housekeeping

offset	description	bits	R/W
0x0	ID
	Reserved	31:4	R
	Design ID 0-prototype0, 1-release1	3:0	R
0x4	DNA part1
	DNA[31:0]	31:0	R
0x8	DNA part2
	Reserved	31:25	R
	DNA[56:32]	24:0	R
0x10	Expansion connector direction P
	Reserved	31:8	R
	Direction for P lines / 1-out / 0-in	7:0	R/W
0x14	Expansion connector direction N
	Reserved	31:8	R
	Direction for N lines / 1-out / 0-in	7:0	R/W
0x18	Expansion connector output P
	Reserved	31:8	R
	P pins output	7:0	R/W
0x1C	Expansion connector output N
	Reserved	31:8	R
	N pins output	7:0	R/W
0x20	Expansion connector input P
	Reserved	31:8	R
	P pins input	7:0	R
0x24	Expansion connector input N
	Reserved	31:8	R
	N pins input	7:0	R
0x30	LED control
	Reserved	31:8	R
	LEDs 7-1	7:1	R/W
	always 0	0	R

## Oscilloscope

offset	description	bits	R/W
0x0	Configuration
	Reserved	31:2	R
	Reset write state machine	1	W
	Start writing data into memory (ARM trigger).	0	W
0x4	Trigger source
	Selects trigger source for data capture. When trigger delay is ended value goes to 0.
	Reserved	31:4	R
	Trigger source: / 1-trig immediately / 2-ch A threshold positive edge / 3-ch A threshold negative edge / 4-ch B threshold positive edge / 5-ch B threshold negative edge / 6-external trigger positive edge - DIO0_P pin / 7-external trigger negative edge / 8-arbitrary wave generator application positive edge / 9-arbitrary wave generator application negative edge	3:0	R/W
0x8	Ch A threshold
	Reserved	31:14	R
	Ch A threshold, makes trigger when ADC value cross this value	13:0	R/W
0xC	Ch B threshold
	Reserved	31:14	R
	Ch B threshold, makes trigger when ADC value cross this value	13:0	R/W
0x10	Delay after trigger
	Number of decimated data after trigger written into memory	31:0	R/W
0x14	Data decimation
	Decimate input data, uses data average
	Reserved	31:17	R
	Data decimation, supports only this values: 1,8, 64,1024,8192,65536. If other value is written data will NOT be correct.	16:0	R/W
0x18	Write pointer - current
	Reserved	31:14	R
	Current write pointer	13:0	R
0x1C	Write pointer - trigger
	Reserved	31:14	R
	Write pointer at time when trigger arrived	13:0	R
0x20	Ch A hysteresis
	Reserved	31:14	R
	Ch A threshold hysteresis. Value must be outside to enable trigger again.	13:0	R/W
0x24	Ch B hysteresis
	Reserved	31:14	R
	Ch B threshold hysteresis. Value must be outside to enable trigger again.	13:0	R/W
0x28	Other
	Reserved	31:1	R
	Enable signal average at decimation	0	R/W
0x30	CH A Equalization filter
	Reserved	31:18	R
	AA Coefficient	17:0	R/W
0x34	CH A Equalization filter
	Reserved	31:25	R
	BB Coefficient	24:0	R/W
0x38	CH A Equalization filter
	Reserved	31:25	R
	KK Coefficient	24:0	R/W
0x3C	CH A Equalization filter
	Reserved	31:25	R
	PP Coefficient	24:0	R/W
0x40	CH B Equalization filter
	Reserved	31:18	R
	AA Coefficient	17:0	R/W
0x44	CH B Equalization filter
	Reserved	31:25	R
	BB Coefficient	24:0	R/W
0x48	CH B Equalization filter
	Reserved	31:25	R
	KK Coefficient	24:0	R/W
0x4C	CH B Equalization filter
	Reserved	31:25	R
	PP Coefficient	24:0	R/W
0x50	CH A AXI lower address
	Starting writing address	31:0	R/W
0x54	CH A AXI upper address
	Address where it jumps to lower	31:0	R/W
0x58	CH A AXI delay after trigger
	Number of decimated data after trigger written into memory	31:0	R/W
0x5C	CH A AXI enable master
	Reserved	31:1	R
	Enable AXI master	0	R/W
0x60	CH A AXI write pointer - trigger
	Write pointer at time when trigger arrived	31:0	R
0x64	CH A AXI write pointer - current
	Current write pointer	31:0	R
0x70	CH B AXI lower address
	Starting writing address	31:0	R/W
0x74	CH B AXI upper address
	Address where it jumps to lower	31:0	R/W
0x78	CH B AXI delay after trigger
	Number of decimated data after trigger written into memory	31:0	R/W
0x7C	CH B AXI enable master
	Reserved	31:1	R
	Enable AXI master	0	R/W
0x80	CH B AXI write pointer - trigger
	Write pointer at time when trigger arrived	31:0	R
0x84	CH B AXI write pointer - current
	Current write pointer	31:0	R
0x90	Trigger debouncer time
	Number of ADC clock periods trigger is disabled after activation / reset value is decimal 62500 or equivalent to 0.5ms	19:0	R/W
0x94	Accumulator control/status
	Reserved	31:2	R
	Accumulator run: / writing an 1 will start accumulating data, writing 0 should not be used, the main configuration register should be used for reset / reading will show the status of the accumulation process (1 – running, 0 - finished)	1	R/W
	Accumulator enable (this disables storing scope data into other buffers)	0	R/W
0x98	Accumulator counter
	on write the number of accumulated samples is specified / on read the status of the incrementing counter is provided for monitoring (0x0 – accumulate 1 sample, 0xffffffff – accumulate 4294967296 samples)	31:0	R/W
0x9C	Accumulator output shift
	Reserved	31:5	R
	specifies by how many LSB bits from the accumulated storage (up to 48 bits) are removed before reaching software, which is limited to 32 bits	4:0	R/W
0xA0	Accumulator data sequence length
	Reserved	31:14	R
	specifies the number of samples accumulated after a trigger, it is limited by the length of the accumulation buffer to 214 = 16384 locations (0x0 – sequence of 1 sample, 0xffffffff – sequence of 16384 samples)	13:0	R/W
0x10000 to 0x1FFFC	Memory data (16k samples)
	Reserved	31:16	R
	Captured data for ch A	15:0	R
0x20000 to 0x2FFFC	Memory data (16k samples)
	Reserved	31:16	R
	Captured data for ch B	15:0	R
0x30000 to 0x3FFFC	Accumulator memory data (16k samples)
	Accumulator data for ch A	31:0	R
0x40000 to 0x4FFFC	Accumulator memory data (16k samples)
	Accumulator data for ch B	31:0	R

## Arbitrary Signal Generator (ASG)

offset	description	bits	R/W
0x0	Configuration
	Reserved	31:25	R
	ch B external gated repetitions	24	R/W
	ch B set output to 0	23	R/W
	ch B SM reset	22	R/W
	Reserved	21	R/W
	ch B SM wrap pointer (if disabled starts at address 0)	20	R/W
	ch B trigger selector: (don't change when SM is active) / 1-trig immediately / 2-external trigger positive edge - DIO0_P pin / 3-external trigger negative edge	19:16	R/W
	Reserved	15:9	R
	ch A external gated bursts	8	R/W
	ch A set output to 0	7	R/W
	ch A SM reset	6	R/W
	Reserved	5	R/W
	ch A SM wrap pointer (if disabled starts at address 0)	4	R/W
	ch A trigger selector: (don't change when SM is active) / 1-trig immediately / 2-external trigger positive edge - DIO0_P pin / 3-external trigger negative edge	3:0	R/W
0x4	Ch A amplitude scale and offset
	out = (data*scale)/0x2000 + offset
	Reserved	31:30	R
	Amplitude offset	29:16	R/W
	Reserved	15:14	R
	Amplitude scale. 0x2000 == multiply by 1. Unsigned	13:0	R/W
0x8	Ch A counter wrap
	Reserved	31:30	R
	Value where counter wraps around. Depends on SM wrap setting. If it is 1 new value is get by wrap, if value is 0 counter goes to offset value. 16 bits for decimals.	29:0	R/W
0xC	Ch A start offset
	Reserved	31:30	R
	Counter start offset. Start offset when trigger arrives. 16 bits for decimals.	29:0	R/W
0x10	Ch A counter step
	Reserved	31:30	R
	Counter step. 16 bits for decimals.	29:0	R/W
0x14	Ch A buffer current read pointer
	Reserved	31:16	R
	Read pointer	15:2	R/W
	Reserved	1:0	R
0x18	Ch A number of read cycles in one burst
	Reserved	31:16	R
	Number of repeats of table readout. 0=infinite	15:0	R/W
0x1C	Ch A number of burst repetitions
	Reserved	31:16	R
	Number of repetitions. 0=disabled	15:0	R/W
0x20	Ch A delay between burst repetitions
	Delay between repetitions. Granularity=1us	31:0	R/W
0x24	Ch B amplitude scale and offset
	out = (data*scale)/0x2000 + offset
	Reserved	31:30	R
	Amplitude offset	29:16	R/W
	Reserved	15:14	R
	Amplitude scale. 0x2000 == multiply by 1. Unsigned	13:0	R/W
0x28	Ch B counter wrap
	Reserved	31:30	R
	Value where counter wraps around. Depends on SM wrap setting. If it is 1 new value is get by wrap, if value is 0 counter goes to offset value. 16 bits for decimals.	29:0	R/W
0x2C	Ch B start offset
	Reserved	31:30	R
	Counter start offset. Start offset when trigger arrives. 16 bits for decimals.	29:0	R/W
0x30	Ch B counter step
	Reserved	31:30	R
	Counter step. 16 bits for decimals.	29:0	R/W
0x34	Ch B buffer current read pointer
	Reserved	31:16	R
	Read pointer	15:2	R/W
	Reserved	1:0	R
0x38	Ch B number of read cycles in one burst
	Reserved	31:16	R
	Number of repeats of table readout. 0=infinite	15:0	R/W
0x3C	Ch B number of burst repetitions
	Reserved	31:16	R
	Number of repetitions. 0=disabled	15:0	R/W
0x40	Ch B delay between burst repetitions
	Delay between repetitions. Granularity=1us	31:0	R/W
0x10000 to 0x1FFFC	Ch A memory data (16k samples)
	Reserved	31:14	R
	ch A data	13:0	R/W
0x20000 to 0x2FFFC	Ch B memory data (16k samples)
	Reserved	31:14	R
	ch B data	13:0	R/W

## PID Controller

offset	description	bits	R/W
0x0	Configuration
	Reserved	31:4	R
	PID22 integrator reset	3	R/W
	PID21 integrator reset	2	R/W
	PID12 integrator reset	1	R/W
	PID11 integrator reset	0	R/W
0x10	PID11 set point
	Reserved	31:14	R
	PID11 set point	13:0	R/W
0x14	PID11 proportional coefficient
	Reserved	31:14	R
	PID11 Kp	13:0	R/W
0x18	PID11 integral coefficient
	Reserved	31:14	R
	PID11 Ki	13:0	R/W
0x1C	PID11 derivative coefficient
	Reserved	31:14	R
	PID11 Kd	13:0	R/W
0x20	PID12 set point
	Reserved	31:14	R
	PID12 set point	13:0	R/W
0x24	PID12 proportional coefficient
	Reserved	31:14	R
	PID12 Kp	13:0	R/W
0x28	PID12 integral coefficient
	Reserved	31:14	R
	PID12 Ki	13:0	R/W
0x2C	PID12 derivative coefficient
	Reserved	31:14	R
	PID12 Kd	13:0	R/W
0x30	PID21 set point
	Reserved	31:14	R
	PID21 set point	13:0	R/W
0x34	PID21 proportional coefficient
	Reserved	31:14	R
	PID21 Kp	13:0	R/W
0x38	PID21 integral coefficient
	Reserved	31:14	R
	PID21 Ki	13:0	R/W
0x3C	PID21 derivative coefficient
	Reserved	31:14	R
	PID21 Kd	13:0	R/W
0x40	PID22 set point
	Reserved	31:14	R
	PID22 set point	13:0	R/W
0x44	PID22 proportional coefficient
	Reserved	31:14	R
	PID22 Kp	13:0	R/W
0x48	PID22 integral coefficient
	Reserved	31:14	R
	PID22 Ki	13:0	R/W
0x4C	PID22 derivative coefficient
	Reserved	31:14	R
	PID22 Kd	13:0	R/W

## Analog Mixed Signals (AMS)

offset	description	bits	R/W
0x0	XADC AIF0
	Reserved	31:12	R
	AIF0 value	11:0	R
0x4	XADC AIF1
	Reserved	31:12	R
	AIF1 value	11:0	R
0x8	XADC AIF2
	Reserved	31:12	R
	AIF2 value	11:0	R
0xC	XADC AIF3
	Reserved	31:12	R
	AIF3 value	11:0	R
0x10	XADC AIF4
	Reserved	31:12	R
	AIF4 value (5V power supply)	11:0	R
0x20	PWM DAC0
	Reserved	31:24	R
	PWM value (100% == 156)	23:16	R/W
	Bit select for PWM repetition which have value PWM+1	15:0	R/W
0x24	PWM DAC1
	Reserved	31:24	R
	PWM value (100% == 156)	23:16	R/W
	Bit select for PWM repetition which have value PWM+1	15:0	R/W
0x28	PWM DAC2
	Reserved	31:24	R
	PWM value (100% == 156)	23:16	R/W
	Bit select for PWM repetition which have value PWM+1	15:0	R/W
0x2C	PWM DAC3
	Reserved	31:24	R
	PWM value (100% == 156)	23:16	R/W
	Bit select for PWM repetition which have value PWM+1	15:0	R/W
0x30	XADC temperature
	Reserved	31:12	R
	Temperature ADC value. / To calculate value into degrees Celsius use this equation: / T(C) = (reg_value * 503.975 / 4095) – 273.15	11:0	R
0x34	XADC VCCPINT
	Reserved	31:12	R
	VCCPINT ADC value. / Value FFFh represents 3V input	11:0	R
0x38	XADC VCCPAUX
	Reserved	31:12	R
	VCCPAUX ADC value. / Value FFFh represents 3V input	11:0	R
0x3C	XADC VCCBRAM
	Reserved	31:12	R
	VCCBRAM ADC value. / Value FFFh represents 3V input	11:0	R
0x40	XADC VCCINT
	Reserved	31:12	R
	VCCINT ADC value. / Value FFFh represents 3V input	11:0	R
0x44	XADC VCCAUX
	Reserved	31:12	R
	VCCAUX ADC value. / Value FFFh represents 3V input	11:0	R
0x48	XADC VCCDDR
	Reserved	31:12	R
	VCCDDR ADC value. / Value FFFh represents 3V input	11:0	R

## Daisy Chain

offset	description	bits	R/W
0x0	Control
	Reserved	31:2	R
	RX enable	1	R/W
	TX enable	0	R/W
0x4	Transmitter data selector
	Custom data	31:16	R/W
	Reserved	15:8	R
	Data source / 0-data is 0 / 1-user data (from logic) / 2-custom data (from this register) / 3-training data (0x00FF) / 4-transmit received data (loop back) / 5-random data (for testing)	3:0	R/W
0x8	Receiver training
	Reserved	31:2	R
	Training successful	1	R
	Enable training	0	R/W
0xC	Received data
	Received data which is different than 0	31:16	R
	Received raw data	15:0	R
0x10	Testing control
	Reserved	31:1	R
	Reset testing counters (error & data)	0	R/W
0x14	Testing error counter
	Error increases if received data is not the same as transmitted testing data	31:0	R
0x18	Testing data counter
	Counter increases when value different as 0 is received	31:0	R

## Power Test

offset	description	bits	R/W
0x0	Control
	Reserved	31:1	R
	Enable module	0	R/W